INSTRUCTION MANUAL

UV RAYS

TYPE RS485

JXBS-3001-UV

VER1.1

WeihaiJXCTElectronicTechnologyCo.,Ltd.

# Ⅰ Brief Introduction

1.1 Product Overview

This product adopts high sensitivity photosensitive probe, the signal is stable and the precision is high. It has the characteristics of wide measurement range, good linear degree, good waterproof performance, convenient use, easy installation and long transmission distance.

1.2 Primary Parameters

| PARAMETERS | TECHNICAL SPECIFICATIONS |
| --- | --- |
| POWER SUPPLY | 12-24V DC |
| POWER | 0.4W |
| OUTPUT SIGNAL | RS485 |
| UV ACCURACY | ±3% (25°C) |
| UV MEASUREMENT RANGE | 0-150W/m2 |
| HUMIDITY RANGE | 0%RH-100%RH |
| TEMPERATURE RANGE | -40℃-80℃ |
| UV LONG-TERM STABILITY | ≤5%/y |
| POWER CONSUMPTION | ≤0.15W(@12V DC , 25℃) |
| WORKING PRESSURE RANGE | 0.9-1.1atm |

1.3 UV Parameters

Figure 1. Influence of different wavelengths on UV intensity

As shown in Figure 1, the wavelength is most accurate around 370nm and the scale factor is 1.

1.4 Temperature and humidity parameters

| PARAMETERS | CONTENT |
| --- | --- |
| HUMIDITY PRECISION | ±3%RH (0%RH-100%RH,25°C) |
| TEMPERATURE PRECISION | ±0.5°C (25°C) |
| HUMIDITY RANGE | 0%RH-100%RH |
| TEMPERATURE RANGE | -25°C-+85°C |
| LONG-TERM STABILITY OF TEMPERATURE | ≤0.1°C/y |
| LONG-TERM STABILITY OF HUMIDITY | ≤1%/y |
| HUMIDITY RANGE | 0%RH-100%RH |
| TEMPERATURE RANGE | -40°C-80°C |
| LONG-TERM STABILITY OF TEMPERATURE | ≤0.1°C/y |
| LONG-TERM STABILITY OF HUMIDITY | ≤1%/y |

1.5 Temperature Parameters

| CONTENT | Minimum | Typical value | Maximum | Unit |
| --- | --- | --- | --- | --- |
| Resolution（14bit） | - | 0.01 | - | ℃ |
| Linear deviation | - | ±0.3 | Figure 2 | ℃ |
| Repeatability | - | ±0.1 | - | ℃ |
| The scope of work | -40 | - | 125 | ℃ |
| Response time (63%) | 5 | - | 30 | second |
| Long-term drift | - | <0.04 | - | ℃/year |

Figure2 Temperature measurement accuracy at different temperatures

As shown in the above table, the maximum deviation is within ±0.5 in the range of 5-60 degrees Celsius, and the deviation increases with less than 0 degrees Celsius and greater than 60 degrees Celsius.

1.6 Humidity Parameter

| CONTENT | Minimum | Typical value | Maximum | Unit |
| --- | --- | --- | --- | --- |
| Resolution（12bit） | - | 0.04 | - | %RH |
| Linear deviation | - | ±3.0 | Figure 3 | %RH |
| Repeatability | - | ±0.1 | - | %RH |
| The scope of work | 0 | - | 100 | %RH |
| Response time (63%) | - | 8 | - | second |
| Long-term drift | - | <0.5 | - | %RH /year |
| Hysteresis | - | ±1 | - | %RH |
| Non-linearity | - | <0.1 | - | %RH |

Figure3 Humidity measurement accuracy under different humidity conditions

As shown in the above table, the deviation increases with a typical deviation of ±3% in the 20-80% RH range, less than 20%, and greater than 80% humidity.

## 1.7 Influence of temperature on humidity measurement

Figure 2 describes the effects of different humidity conditions on the humidity measurement. The following table describes the effects of different temperatures on the humidity measurement accuracy.

Figure 4 Relative humidity at different temperatures

As shown in the above table, in the range of 15 to 55 degrees Celsius and 30 to 80 degrees of humidity, the accuracy of humidity is the highest, which is ± 4.5%. In other cases, the humidity increases.

1.8 System frame Diagram

FIGURE 5   SINGLE-ENDED

FIGURE 6   MUTIPLE-ENDED

# Ⅱ Hardware Connections

2.1 CHECKING BEFORE INSTALLATION

Check the list of devices before installation:

TABLE 1   List of Devices

| Name | Number |
| --- | --- |
| THE SENSOR DEVICE | 1 |
| 12V POWER ADAPTER（Optional） | 1 |
| WARRANTY CARD / CERTIFICATE | 1 |
| THE USB TO 485 DEVICE（Optional） | 1 |

2.2 Interface Description

The power interface is wide voltage input 12-24V. When connecting the 485 signal cable, note that the A/B lines cannot be reversed. The addresses of multiple devices on the bus cannot conflict.

FIGURE 7   PHYSICAL PICTURE

TABLE 2   Wiring Sequence

|  | Line Color |  | Description |
| --- | --- | --- | --- |
| Power | Brown |  | Power supply Positive ( 12-24V DC ) |
|  | Black |  | Power supply Negative |
| Communication | Yellow |  | 485-A |
|  | Blue |  | 485-B |

Precautions: Please take care not to take the wrong line sequence, the wrong wiring will cause the equipment to burn.

The factory default is to provide 0.6 meters long wire rods. Customers can extend the wire rods as required or sequentially.

Note that there are no yellow lines in the wire sequences that may be provided in some factory batches, in which case the gray line is equivalent to replacing the yellow line.

2.3 Installation

Ultraviolet transmitters are greatly affected by the angle of incidence, so the following two mounting options are recommended:

1. Install horizontally with the ground plane so that the sensory position is facing the sky. All-day measurements can be made as a fixed location.

2. If you want to measure the current real-time UV intensity, you can make the photo-sensing position directly opposite to the incident position of sunlight.

# Ⅲ Configuration Tool Installation and use

We provide CONFIGURATION TOOL , which can be easily used to test our sensor device.

3.1 Sensor Access Computer

Transmitter can be connected to PC with the RS485 to USB adapter. You can check the COM port number through Device Manager (right click My Computer).

3.2 How To Use Configuration Tool

Please note that this software can only test one device at the same time. After connecting the physical device, click the CONNECT button to read the information. In the UNCONNET state, you can modify BAUD and ADDR in COMMUNICATION SETTINGS.

Under the software, different check boxes can be selected according to different situations. For example, you can choose the GAS option to test the RS485 OXYGEN SENSOR , you can choose the NORMAL option to test the RS485 TEMPERATURE AND HUMIDITY SENSOR .

# Ⅳ Communication Protocol

4.1 Communication Basic Parameters

TABLE 3  Communication Basic Parameters

| PARAMETERS | CONTENT |
| --- | --- |
| Protocol | Modbus RTU |
| Data bits | 8 bit |
| Parity bit | No |
| Stop bit | 1 bit |
| Error checking | CRC (redundant loop code) |
| Baud rate | 2400 bps/ 4800 bps/ 9600 bps can be set factory defaults to 9600 bps |

For more information about MODBUS RTU please visit the website"www.modbus.org ".

4.2 Register Address

| Register Address | Plc Configuration Address | Content | Operation |
| --- | --- | --- | --- |
| 0000H | 40001 | Humidity(unit 0.1%RH) | Read-Only |
| 0001H | 40002 | Temperature(unit 0.1℃) | Read-Only |
| 0008H | 40009 | UV Rays（unit0.1W/㎡） | Read-Only |
| 0100H | 40101 | Device Address (0-252) | R/W |
| 0101H | 40102 | Baud Rate (2400/4800/9600) | R/W |

TABLE 4  Register Address

4.3 Communication example

4.3.1 Read Device Address 0x01's UV Value

TABLE 5   Inquiry Frame

| Address Code | Function Code | Start Address | Data Length | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x00 0x08 | 0x00 0x01 | 0x05 | 0xC8 |

TABLE 6   Answer Frames

( For example, the reading is 68.0W/㎡ )

| Address Code | Function Code | Returns to The Number Of Valid Bytes | Carbon Dioxide Value | Check Digit Low | Check Digit High |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x02 | 0x02 0xA8 | 0xB8 | 0x9A |

UV：

02A8 H(hexadecimal )=680=>UV=68.0W/㎡

4.3.2 Read Device Address 0x01's Temperature And Humidity Value

TABLE 7   Inquiry Frame

| Address Code | Function Code | Start Address | Data Length | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x00,0x00 | 0x00,0x02 | 0xC4 | 0x0B |

TABLE 8   Answer Frames

| `Address Code | Function Code | Number Of Valid Bytes | Humidity Value | Temperature Value | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x04 | 0x02 0x92 | 0xFF 0x9B | 0x5A | 0x3D |

Temperature：

FF9BH ( hexadecimal ) =-101=> temperature =-10.1℃

Humidity:

292H( hexadecimal ) =658=>humidity =65.8%RH

4.3.3 Read Device Address 0x01's UV,Temperature And Humidity Value

TABLE 9   Inquiry Frame

| Address Code | Function Code | Start Address | Data Length | CRC_L | CRC_H |
| --- | --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x00,0x00 | 0x00,0x09 | 0x85 | 0xCC |

TABLE 10  Answer Frame

| Address Code | Function Code | Number Of Bytes | Humidity Value | Temperature Value |
| --- | --- | --- | --- | --- |
| 0x01 | 0x03 | 0x12 | 0x02 0x92 | 0x01 0x1C |
| 12 Useless Bytes | UV value | Check code |
| 00 | 0x02 0xA8 | 0x33 0x4B |

UV：

02A8 H(hexadecimal)=680=>UV=68.0W/m2

Temperature：

011C H(hexadecimal)=284=>Temperature=28.4℃

Humidity：

0292 H(hexadecimal)=658=>Humidity=65.8%RH